Taprisiot Application User Manual

| Abstract | This documents is the user manual of the ConoscopeDemo application |
| --- | --- |
| Version | 0.4 |
| Status | Draft |
| Date | 2020/11/18 |

Revision history

| Version | Date | Content |
| --- | --- | --- |
| 0.1 | 2020/06/03 | Initial version |
| 0.2 | 2020/06/25 | Add “log file” chapter |
| 0.3 | 2020/07/01 | Add export file naming options Add ROI Add AutoExposure Measurement Area Add AutoExposure level and exposure time thresholds Add CaptureSequence SaveCapture option Add ConvertRaw feature |
|  | 2020/07/02 | Add AE exposure time granularity |
|  | 2020/07/06 | Add <AeExpoGran> tag in ExportFormat name |
| 0.4 | 2020/11/18 | Add streaming options |

|  | ELDIM 1185 Rue d’Epron (Ancienne) 14200 Hérouville Saint-Clair France |
| --- | --- |
|  | Copyright © 2020 All rights reserved. Printed in France. |
|  | ELDIM, the ELDIM logo and other product names referenced herein are trademarks of ELDIM. Other product names, designations, logos, and symbols are trademarks or registered trademarks of their respective owners. |
|  | NO WARRANTY. The technical documentation is being delivered to you AS-IS and ELDIM makes no warranty as to its accuracy or use. Any use of the technical documentation or the information contained therein is at the risk of the user. Documentation may include technical or other inaccuracies or typographical errors. ELDIM reserves the right to make changes without prior notice. |
|  | ELDIM considers information included in this documentation as Confidential Information. The access and use of this confidential information are subject to the terms and conditions of the Software license agreement, with which you agree to comply. This documentation cannot be reproduced in any way without the prior agreement and written permission of ELDIM. |

Table of contents

1	Introduction	6

2	Install	6

2.1	Requirement	6

2.2	Application	6

2.3	User interface	7

3	Application version	7

4	Calibration Data	8

4.1	Manually	9

4.2	Download Calibration data from Taprisiot device	10

4.2.1	open	10

4.2.2	Read data	10

4.2.3	Read Status	10

5	Capturing an image	11

5.1	Configuring capture folder	11

5.2	Open the device	11

5.3	Setup	11

5.4	Image Capture	11

5.4.1	Measure command	11

5.4.2	MeasureAE command	11

5.4.3	Measure output	13

5.5	Export Capture	14

5.5.1	Export Capture options:	14

5.5.2	Raw Data Format	15

5.5.3	Processed Data Format	16

5.6	Display Capture	16

5.7	Next Step	16

6	Closing the device / Reset	17

7	Streaming	17

7.1	Usage	17

7.2	Display options:	17

8	Capture Sequence	18

8.1	Description	18

8.2	Command	18

8.3	Result	20

9	Processing captured data	21

9.1	Processing RawData	21

9.2	Processing CaptureSequence	21

10	Log file	22

10.1	Introduction	22

10.2	Configuration	22

10.3	Log file	22

# Introduction

Conoscope Application drives the taprisiot device.

The application handles following features:

- Calibration data

- Taprisiot configuration: temperature and filter selection

- Capture: with specific exposure time or auto exposure option

- Export data: Raw Data or Processed data

- Capture Sequence: Capture required images to generate X, Y, Z images

# Install

## Requirement

Euresys drivers to drive the Frame Grabber: coaxlink-win10-x86_64-11.0.3.82.exe
Once install, launch GenICam browser. The application may request to update frame grabber firmware.
Please select Coaxlink Quad G2 (2 camera), update firmware. Once done shut down and unplug the PC to be sure the frame grabber board is properly switched off.

## Application

The application is ConoscopeDemo.exe. The folder contains all required dependencies.

## User interface

Warning:
When launching the application, it might be possible that an error pops-up indicating that a dll is missing.
Please install vcredist_x64.exe and/or vc_redist.x64.exe

# Application version

The application version and libraries version are indicated on the interface:

# Calibration Data

To run properly, the application must access calibration data of the Taprisiot.
The folder where the calibration data are stored can be modified using SetConfig command:
set path in cfgPath and press SetConfig

Note: It is only possible to change the Calibration data folder after application start-up. Once Taprisiot is open, the folder can not be changed any more.

Calibration Data folder is populated by 2 ways:

- Manually. The user has calibration data and can copy it in the appropriate folder

- From Taprisiot: Data are stored in the memory of the Taprisiot and can be retrieved.

## Manually

Copy calibration data into the folder. 
remark: the subfolder matches with the camera SN (not the serial number of the taprisiot) 
To know the camera SN, press Open command. The Camera SN is displayed in the log.

## Download Calibration data from Taprisiot device

calibration data is stored in the camera. It is possible to populate the folder with those values.
Remark: This upload takes a lot of time (around 30 minutes)

### open

### Read data

Launch read using CfgFileRead command

### Read Status

Then regularly check Cfg with CfgFileStatus command until the process is complete.

# Capturing an image

## Configuring capture folder

When an image is captured, it is stored in a specific folder.
Change the path in capturePath and press SetConfig

## Open the device

Initially Taprisiot is in Idle state.
Press Open to bring the device in opened state.

## Setup

Once opened, it is possible to setup the device

- Targeted temperature: 	the device will be regulated at this temperature

- Filter: 			BK7, Mirror, X, Xz, Ya, Yb, Z, IrCut

- Nd filter: 		density 0, 1, 2, 3, 4

- Iris:			aperture: 2mm, 3mm, 4mm, 5mm. Aperture must be changed manually.

Press Setup to go in ready state

Current configuration can be retrieved with SetupStatus command

## Image Capture

### Measure command

Measure command will capture the image at the specified ExposureTime

### MeasureAE command

MeasureAE will capture the image after an AutoExposure sequence.

During MeasureAE processing, it is possible to know the status of the processing (MeasureAEStatus) and to stop the processing (MeasureAECancel)

MeasureAE command works with several iterations and will take more time than Measure command.

During the algorithm, an image is captured and processed.

The brightest pixel of the measurement area is measured, and the exposure time is tuned so this value matches a specified level of saturation. By default, the exposure time is tuned so the return value is 80% of saturation.

Measure AE Parameters:

| Measurement Area | Defines the area where the measurement is done. note: the bigger the area is, the slower the algorithm is. |
| --- | --- |
|  | AEMeasAreaHeight [0, 6004] AEMeasAreaWidth [0, 7920] AEMeasAreaX [0, 7920] AEMeasAreaY [0, 6004] Set All parameters to 0 to disable |
| Level | Defines the target of the algorithm. If Level is set to 80%, the exposure time is calculated so the brightest pixel of the measurement area is 80% of the saturation |
|  | AELevel (percent) |
| Threshold | Boundary in the exposure time (in micro seconds) |
|  | AEMinExpoTimeUs AEMaxExpoTimeUs |
| Granularity | Defines the granularity of the Exposure Time calculated by the algorithm. AEExpoTimeGranularityUs |

### Measure output

Measure Commands hold data internally.
To store capture, use command ExportRaw or ExportProcessed.
To display capture, use DisplayRaw or DisplayProcessed.

## Export Capture

From state MeasureDone, use Commands ExportRaw and ExportProcessed to store capture in binary format. Those commands do not perform a capture (Measure commands does, see previous chapter).

### Export Capture options:

| capturePath | Folder where capture are exported |
| --- | --- |
| ROI | Defines the Region of interest of the exported image. (this apply only on processed capture) 0 XLeft XRight 0 +---+------------+---+ ¦ ¦ YTop + ++++++++++++++ ¦ ¦ ++++++++++++++ ¦ ¦ ++++++++++++++ ¦ ¦ ++++++++++++++ ¦ YBot + ++++++++++++++ ¦ ¦ ¦ +--------------------+ Each value in range [0, 6001] |
| fileNamePrepend | Add a string the beginning of the exported file name |
| fileNameAppend | Append a string at the end of the exported file name |
| ExportFormat | Specify the format of the exported file name |
|  | <TimeStamp> <Filter> <Nd> <Iris> <ExpoTime> <NbAcq> <Height> <Width> <SatFlag> <SatLevel> <AeExpoGran> |
|  | default format <TimeStamp>_filt_<Filter>_nd_<Nd>_iris_<Iris> |

### Raw Data Format

A capture is .bin file associated to a .json file containing information about the configuration of the device.
A binary file can be opened using ImageJ

File > Import > Raw 
         Image type :	16 bits unsigned
         Width :		7920
         Height:		6004
         Little endian byte order

### Processed Data Format

Image processing:
        Bias compensation
        Sensor defect correction
        Sensor PRNU correction
        Linearisation (depends on Setup configuration)
        FlatField (depends on Setup configuration)

File > Import > Raw 
         Image type :	16 bits signed
         Width :		6001
         Height:		6001
         Little endian byte order

## Display Capture

DisplayRaw and DisplayProcessed allow to display the image captured with Measure command

Those commands do not perform a capture

## Next Step

Once a capture has been done, it is possible Measure again or to Setup and Measure

# Closing the device / Reset

In any state, the device can be closed or reset.
Reset command will perform a power sequence on the Taprisiot (it takes about 2 minutes)

# Streaming

## Usage

When the device is in Ready state (after a Setup command). It is possible to display stream by pressing DisplayStream.

Streaming options: 	AutoExposure: 	AE parameters are set with SetOption command 
			bProcessed: 	Capture data are processed. 
					By default, only defect pixels are removed.
					Configuration can be changed in DisplayStreamOption.json file

## Display options:

Zoom		use the mouse wheel. 
Cursors		right click on the image and select the option

# Capture Sequence

## Description

Capture sequence captures necessary images to generate X, Y, Z images.
5 capture are used:
	- filter X
	- filter Xz
	- filter Ya
	- filter Yb
	- filter Z

## Command

Command CaptureSequence can be called from Opened state.

Parameters:
	sensorTemperature: 	Target temperature for the Taprisiot
	waitForTemp: 		The sequence waits till the target temperature is reached
	nd: 			Setup the ND filter value
	iris:			Indicate the iris installed on the Taprisiot
	exposure time:		option 1: 	manual exposure, all captures are done 
						with the same exposure time
					
				option 2: 	exposure time of each capture is defined in 
						CaptureSequenceExposureTime.json
					
				option 3: 	AutoExposure

CaptureSequenceExposureTime.json the exposure time for each filter and is in the application directory

{
		"ExposureTimeUs": {
			"Filter_X": 10000,
			"Filter_Xz": 10000,
			"Filter_Ya": 10000,
			"Filter_Yb": 10000,
			"Filter_Z": 10000
		}
	}

Once the processing is started, it is possible to cancel it with CaptureSequenceCancel command

## Result

Once the process is done, X, Y, Z images can be found in capture folder (SetConfig)

Note: If bSaveCaptures is checked, all intermediate captures are stored. Then, they can be processed latter.

Images can be opened with ImageJ

# Processing captured data

It is possible to process raw data already captured.
From Idle state, set emulatedCamera option in SetConfig command (press SetConfig button to set)

## Processing RawData

Follow the steps described in chapter “Capturing an image”.
Parameter for Measure command is the path of the bin capture to process.
(note the associated json file must be present)

Then ExportProcessed command will generate the processed image using the calibration data.

## Processing CaptureSequence

Capture paths to be processed are defined in CaptureSequenceCaptures.json (in the application folder)

{
		"FilePath": {
			"Filter_X": "capture_X.bin",
			"Filter_Xz": "capture_Xz.bin",
			"Filter_Ya": "capture_Ya.bin",
			"Filter_Yb": "capture_Yb.bin",
			"Filter_Z": "capture_Z.bin"
		}
	}

Remark: by default, CaptureSequenceCaptures.json does not exists. When CaptureSequence is lanched without this file, it will generate a sample file. Then it is possible to update it.

To process the data, follow the steps described in chapter “Capture Sequence”

# Log file

## Introduction

Log information can be saved in file.

## Configuration

In installation folder:

Edit demo.json file and add LogMasks:

{
    "Application": {
        "LogMasks": [
            "State",
            "StateMachine",
            "Worker"
        ],

## Log file

Log files are:

LogConoscope.txt
LogPipeline.txt